Paralysis Hand Recovery Machine – System Plan

0) Quick Summary

A 3-phase rehab device (exoskeleton glove + forearm/wrist rig) that supports:

Phase 1 (Passive/Robot-driven): Safely moves the patient’s hand in three modes - finger flex/extend (fist), wrist ulnar/radial deviation (left/right), wrist flex/extend (forward/backward).

Phase 2 (Active/Patient-driven): Measures the patient’s voluntary capability in the same three DoFs without assistance (or with minimal assist-as-needed).

Phase 3 (Analytics & ML): Logs time, angle, torque/force, repetitions, and pain/effort ratings; visualizes progress with dynamic report generation using Power BI/Tableau or use Blynk; estimates recovery time using a simple and explainable ML model, upgraded later to more advanced models.

1) Clinical & Safety Notes (non-negotiables)

- Therapist-approved ROM limits: Per-patient configurable limits for finger flex/extend, wrist flex/extend, and ulnar/radial deviation. Hard mechanical stops + software limits.

- Torque/Force ceilings: Current limiting on motors + real-time torque/force monitoring. Abort if exceeding threshold for >100 ms.

- Emergency stop: Big red e-stop on device + software e-stop in UI. Power stage de-energizes, brakes engage (if available).

- Smoothness: Jerk-limited trajectories (S-curves) to avoid sudden pulls.

- Comfort & hygiene: Soft liners, adjustable straps, easily cleanable surfaces.

- Data privacy: Local encryption at rest; export with patient consent.

- Regulatory pathway (future): Design toward IEC 60601-1 (electrical safety), 60601-1-2 (EMC), risk mgmt per ISO 14971, usability per IEC 62366, basic software lifecycle per IEC 62304.

- Clinical validation: Start with IRB/ethics approval and small feasibility study before claiming clinical efficacy.

1a) Initial Calibration & Setup (First Use Only)

- Manual Calibration Routine: On first setup, the therapist/caregiver manually moves the patient’s hand to their maximum safe range in each DoF:

  - Fully open → fully closed fist (finger flex/extend).

  - Wrist forward → wrist backward (flex/extend).

  - Wrist left → right (ulnar/radial deviation).

- Recording Limits: The system records encoder/IMU angles at each extreme and saves them as per-patient software limits.

- Safety Margins: A 5–10° margin is automatically applied to avoid overstretching.

- Persistent Storage: Limits are stored in patient profile (JSON/EEPROM/DB) and automatically loaded in every session.

- Re-calibration: Can be repeated at any time if patient’s ROM improves or straps/device are repositioned.

2) System Architecture

Mechanical

- Finger actuation (1–5 DoFs depending on budget): tendon-driven cable system per finger or a single flexion/extension bar for all fingers (MCP-dominant).

- Wrist module:

  - Flexion/extension hinge with adjustable axis alignment to wrist joint.

  - Ulnar/radial deviation joint (orthogonal axis).

  - Quick-release mounts and soft straps.

Electronics

- Controller: STM32, ESP32-S3, or similar for low-latency loops; optional single-board computer (Raspberry Pi/Jetson) for UI + ML.

- Drivers: Current-controlled BLDC/servo drivers (e.g., ODrive/MadgeTech alternatives) or high-quality hobby servos with feedback for PoC.

- Sensors:

  - Encoders on each joint/tendon spool (≥12-bit).

  - Force/torque: In-line load cell or FSRs in glove fingertips + wrist strap; optional 6-axis force-torque for advanced control.

  - IMU on hand backplate to sanity-check angles.

  - EMG for bio-intent detection.

User Controls & Feedback:

- Potentiometer (speed control): Adjustable knob to scale exercise speed (rep cadence, trajectory velocity). Read via ADC on MCU.

- LCD Display (status feedback): Small LCD/OLED (e.g., 20×4 character LCD or 1.3" OLED) to show:

  - Current exercise/mode (Passive / Active / Assessment)

  - Repetition count & speed level

  - Error/safety warnings (e.g., torque exceeded, E-stop triggered)

  - Patient progress (ROM reached vs. target in real time)

Software

- Low-level control: 1 kHz inner loop (position/torque) on MCU; jerk-limited trajectory generator.

- Mid-level modes: Passive, Active-assist, and Assessment routines.

- High-level UI & logging: Desktop/web app (Python + FastAPI + Qt/Web) to set protocols, visualize, and export CSV/Parquet for BI tools.

3) Degrees of Freedom (DoFs) & Modes

DoFs covered:

- Finger flexion/extension (make/relax fist) — single coupled DoF initially.

- Wrist ulnar/radial deviation (left/right).

- Wrist flexion/extension (forward/backward).

Modes

- Phase 1: Passive (robot-driven)

  - Predefined ROM and cadence (e.g., 10 reps × 3 sets).

  - Impedance control (stiffness + damping) to remain compliant while tracking target trajectory.

- Phase 2: Assessment (patient-driven)

  - No active torque (or minimal gravity compensation).

  - Measures max angle, time-to-target, smoothness (jerk), and tremor power.

  - Active-assist variant: If patient stalls > T seconds below threshold force, provide minimal assist.

- Phase 3: Analytics & ML

  - Aggregates per-session metrics; estimates time-to-target recovery levels with survival regression or learning-curve models.

4) Control Strategy (concise)

Trajectory generation:

- S-curve (jerk-limited) profiles between angle waypoints; per-patient ROM limits.

Impedance control (Phase 1):

- τ_cmd = K(θ_ref − θ_meas) + D(ω_ref − ω_meas) + τ_ff (gravity/friction).

- Tune K/D conservatively; watchdog for high τ or error.

Assessment logic (Phase 2):

- Cue target (angle or %ROM).

- Start timer; record peak angle, time to reach 80% target, path variability, tremor index (RMS of 4–12 Hz band).

- If assist-as-needed: add τ_assist only if dθ/dt < ε for > T_stall, capped at τ_max.

5) Key Metrics to Log

- Kinematics: angles (°), angular velocity, jerk; ROM used (%).

- Kinetics: commanded torque/current (proxy), measured force/torque (N/N·m).

- Performance: time-to-target (s), success rate (% reps completed), repetitions, hold time.

- Quality: movement smoothness (spectral arc length), tremor power (RMS 4–12 Hz), path deviation.

- Subjective: pain (0–10), effort (Borg scale), fatigue flags.

- Session metadata: patient ID (pseudonymized), therapist ID, device config, firmware version.

6) Data Schema (for CSV/DB)

patients

- patient_id, sex, age, side (L/R), diagnosis, baseline_date

sessions

- session_id, patient_id, date, therapist_id, mode (passive/assessment/assist), notes

trials (one per DoF/target in a session)

- trial_id, session_id, dof (fingers/wrist_fe/wrist_ud), target_angle_deg, target_type (abs/%ROM),

- start_time, end_time, success (bool), time_to_80pct_s, max_angle_deg, assist_applied (bool), pain_0_10

signals (optional large table)

- trial_id, t_ms, angle_deg, velocity_dps, torque_cmd_nm, force_n, imu_quat_wxyz

device_config

- config_id, fw_version, K_impedance, D_impedance, max_current, rom_limits

7) ML Approach for Recovery Time Estimation

Goal: Predict time to reach predefined functional milestone (e.g., 80% baseline ROM, or therapist-set goal) using early-session data.

Baselines (start simple):

- Linear regression / Elastic Net on features from first N sessions → predict remaining sessions to goal.

- Gradient Boosting (XGBoost/LightGBM) for nonlinearities and interactions.

- Survival analysis if censoring exists (patient drops out): Cox proportional hazards or GBM-based survival models.

Example features (per patient up to session k):

- Trend of time-to-target (slope), best ROM so far, last-session smoothness, tremor index change, average assist torque, session adherence.

- Demographics & diagnosis (if available and approved).

Model hygiene:

- Train/validation split by patient (no leakage).

- Report MAE for “sessions remaining” and calibration.

- Keep explainability (feature importances, SHAP) for clinician trust.

8) Visualization & Reporting

- Live session dashboard: per-rep angle/torque traces, safety flags, and rep counter.

- Progress report: weekly chart of ROM, time-to-target, smoothness, assist %.

- Outcome estimate: funnel chart with predicted recovery window and confidence band.

Power BI / Tableau data model:

- Import patients, sessions, trials (and optionally signals).

- Create measures: Avg Time-to-80%, Best ROM, %Sessions with Assist, Trend lines.

- Precompute tidy CSVs after each session for plug-and-play refresh.

9) Bill of Materials (MVP options)

- Mechanical: 3D-printed wrist module + aluminum hinges; Bowden cables for finger actuation; soft glove.

- Actuators:

  - Wrist FE/UD: Small BLDC gimbal motors with planetary gearboxes (ratio 20–50:1) + magnetic encoders.

  - Fingers: High-torque servo (e.g., Dynamixel X-series) or BLDC with tendon spool.

- Drivers: Dual-channel BLDC/servo drivers with current sensing; emergency cutoff relay.

- Sensors: Inline load cell (0–50 N), FSRs for fingertip contact, 6-DoF IMU, 12–14-bit encoders.

- Compute: ESP32-S3 (control) + Raspberry Pi 4/5 (UI/ML) or single Jetson Nano/Orin Nano.

- HMI: 7–10" touchscreen or laptop; physical e-stop.

10) Firmware/Software – Pseudocode Snippets

Impedance-controlled passive rep (Phase 1):

θ_target(t) = S_curve_profile(start=θ0, end=θ1, T=rep_time)

loop at 1 kHz:

θ, ω = read_encoder()

τ_limit = torque_ceiling(dof)

τ_cmd = clamp(K*(θ_target-θ) + D*(ω_target-ω) + τ_ff, ±τ_limit)

if force_sensor > F_max or |τ_cmd| near τ_limit for >100 ms: abort_to_safe()

write_motor_current(τ_cmd)

log(t, θ, ω, τ_cmd, force)

Assessment trial (Phase 2):

cue_user(target_angle)

start_timer()

while time < T_max and angle < 0.8*target:

angle, vel = read_encoder()

if assist_mode and stalled(vel, time):

τ_cmd = min(τ_cmd + Δ, τ_assist_max)

else:

τ_cmd = 0

log(t, angle, vel, τ_cmd)

end

metrics = {time_to_80pct, max_angle, tremor_power, smoothness}

Data pipeline (Phase 3):

after each session:

compute per-trial metrics → write trials.csv

append sessions.csv summary

export tidy dataset for BI and for ML training

11) Acceptance Tests (MVP)

- Safety cut-off: Apply unexpected resistance; device must cut power < 100 ms.

- ROM accuracy: Command +30° FE → measured within ±2°.

- Repeatability: 100 passive reps: std dev of peak angle < 2°.

- Assessment reliability: Time-to-80% repeatability CV < 10% in healthy subject.

- Data integrity: No missing timestamps; CSV schema validates.